SolOhm – Monitor	7/1/2015
High-Precision Panel/String Monitor

Feature Set
Use Cases Supported	● Single panel power verification with temperatures / ● Microinverter MPPT verification and sweep capture / ● String power verification with temperatures / ● String inverter MPPT verification and sweep capture / ● Electronic load sweep capture
Tri-range Voltmeter LTC2348-18	Full scale Resolution / 60VDC < 1mV / 120VDC < 2mV / 700VDC < 10mV
Ammeter LTC2348-18	10A < 0.2mA
Temperature Sensor Inputs LTC2983	5 RTD – 2,3,4 wire / 10 Thermocouples - B, E, J, K, N, S, R, T / 0.1°C accuracy and 0.001°C resolution
Operating Temperature Range	-20°C to 60°C, holds measurement accuracy over range
Real Time Clock	NTP time reference synchronization / Disconnected accuracy < 0.3sec/day
Power Supply Turnigy Lipo	Rechargeable/replaceable COTS 3-cell LiPo battery / Parasitically panel powered, < 0.01% consumption / Consumption disabled during data acquisition
Sampling	Simultaneous voltage and current measurement / No multi-channel multiplexer / 18-bit, 200K samples/sec max, full differential / Configurable sample rate between 1/5min – 200K/sec
Data Storage uSDHC	4GB non-removable – FAT32 file system, timestamped data, date-based filenames, CSV format
Configuration	Bluetooth enabled cell phone application (IOS/Android) / Settings: user label, sample rate, data transfer frequency/destination URL, filename format, WiFi configuration, timezone, calibration constants
Communication	Bluetooth Low Energy configuration and data display, phone app will show live summary table of all field devices, WiFi data transfer
Calibration	Requires power supply and precision voltmeter/ammeter
Enclosure	3”x5”x3”, plastic, IP65, MC4 pigtail connections, mountable, no high voltage access, no buttons, no display, safety markings, viewable serial number
Price	$990.00